## Slide 1: Midterm Project Report Back to the Basics – Forward to the Future
for
IoT based Smart Manufacturing
Ganesh Kumar (MSM17B034)
Dhilipan SP (MSM17B002)
Chandralekha R (MSM17B027)
On
Development of IoT powered Low cost Vibration Spectrum Analyzer

## Slide 2: In a Big View !

## Slide 3: Problem Statement:
Modern technological processes require the continuous monitoring of many parameters of technological equipment. The most important parameters are mechanical parameters such as periodic mechanical vibrations. Vibration monitoring systems make possible to solve a lot of the problems coupled with vibrations that appear during the technological equipment operations. Thus, DEVELOPMENT OF IoT BASED VIBRATION MONITORING AND SPECTRUM ANALYSIS SYSTEMS FOR TECHNICAL OBJECTS is an important task.

## Slide 4: Our Objective…
To design a low cost alternative to vibration monitor and spectrum analyzer using the Raspberry Pi microcomputer and 3-axis digital MEMS ADXL345 accelerometer.
To study if inexpensive vibration monitoring systems could be suitable for condition monitoring
To highlight the impact of different components of the signal chain to the measured vibration signal itself and familiarize the reader with the signal chain found in vibration monitoring.
To analyze vibration parameters in order to predict and prevent possible accidents, thus reducing the costs associated with the failure of the cutting tools, expensive parts and assemblies of the CNC machine

## Slide 5: Progress
Developed a Software UI for the Vibrational analysis
  Application Specification (MATLAB / LabVIEW)
  Target Platform Specification (Remote/ On-board)
  Binding
  HW and SW Estimation
  Partitioning

## Slide 6: Progress
Studied the previous work done in this area ( IEEE’s Cited research papers )
Familiarized ourselves with the MATLAB Development Environment
Understood the overall framework of the project
The role of FFT analysis in the project

## Slide 7: Progress
Implemented FFTplot.py
Worked in Linux environment but failed in raspberry
WHY? Version constraints 2.7  3.6
Matpy lib contradicts……..
Did in Mathscript in Matlab
Succeded in LabVIEW by creating FFT.VI
Enclosed in Drive all the .py, .m, .vi files

## Slide 8: FFT plot UI created in LabVIEW

## Slide 9: FFT plot UI created in Py

## Slide 10: Back to the Basics…
Vibration
  Simple Harmonic Motion
    Oscillation about a Reference Point
    Modeled Mathematically as…
Literature Review

## Slide 11: Back to the Basics…
Period, T
RMS
Unit Circle
Peak-to-Peak
0 to Peak
0
Fig.1 Frequency Cycle

## Slide 12: Basic Signal Attributes
  Static
    Slowly Changing
    Temperature
  Dynamic
    Sensor must respond in fractions of a Second
    Vibration, Amperage, Pressure

## Slide 13: Back to the Basics…
Dynamic Signal Fundamentals
  Amplitude
  Frequency
  Timing
  Shape
Amplitude
  Proportional to the severity of vibratory motion
  Expressed as
    Peak to Peak
    Zero to Peak
    RMS
Frequency
  Determined by the reciprocal of the Period
    CPS or Hz
    RPM
    Orders
Timing, or Phase
  Represented by the time delay between two signals
  Leading
  Lagging
Signal Shape
  Waveform
    Simple
    Complex
    Pattern Recognition

## Slide 14: Peak and RMS Comparison
Fig.2 RMS comparison

## Slide 15: Relationships of Acceleration, Velocity and Displacement
Fig.3 Relationship between v, a, d

## Slide 16: The Big Picture
Sensor(s)
Cables
Signal Conditioning
Data Acquisition & Storage
Communications
Remote Analysis and Diagnostics

## Slide 17: Acceleration Sensors
Pro’s and Con’s
  Pro’s
    Measures Accel.
    Small Size
    Easily Installed
    Large Frequency Range (1-10,000 Hz)
  Con’s
    Measures Acceleration (requires Integration to Vel.)
    Susceptible to Shock & Requires Power
Fig.4 Types of acc. meter

## Slide 18: Voltage or Current?
Current Output Accelerometers
  4-20 mA Output
    Proportional to Dynamic Signal and/or Overall
Voltage Output Accelerometers
  Preferred in U.S.
  Generally 100mV per g Sensitivity

## Slide 19: AC and DC Signal Components
Signals have both AC and DC
  AC considered the “Dynamic” Signal
  DC is the “Static” Signal
    Displacement Probes – Set “Gap” for DC
    Accelerometers – “Bias” voltage is DC

## Slide 20: AC and DC Signal Components
How AC and DC work together:
  AC signal “rides” the DC bias (VB)
    Affects the Dynamic Range of the Sensor.

## Slide 21: Power Circuit for Accelerometers
“Strips off” DC Voltage

## Slide 22: Grounds
A Potential Problem Source
  Ground Loops
    Caused when two or more grounds are at different potentials
    Sensors should be grounded only at the sensor, not the monitoring rack!

## Slide 23: Sensor Cables
Coaxial with BNC Connectors
  Long Coaxial can become antennas!
Twisted, Shielded Pair
  Teflon Shield – ground at only one end!

## Slide 24: Sensor Cables
Driving Long Cables
  Under 90 feet, cable capacitance no problem –Cable Capacitance spec’d in Pico-farads per foot of cable length
  Over 90 feet or so, CCD must supply enough current to charge the cable as well as the sensor amplifier.
    May result in amplifier output voltage becoming “Slew Rate Limited”

## Slide 25: Sensor Cables
Output of Sinusoid looks like this:
What’s Happening?
  The + part of the signal is being limited by the current available to drive the cable capacitance.
  In the – part of the sin wave, the op-amp must “sink” the current being discharged by the cable capacitance.

## Slide 26: Sensor Cables
Practical Effect:
  Signal distortion produces harmonics
  May lead to vibration signals being misinterpreted.
  To calculate the maximum frequency for a length of cable:

## Slide 27: Methodology

## Slide 28
Signal Processing Flow
FFT
Waveform
Spectrum
Transducer
Amplitude
Amplitude
Time
Frequency
Data Collector/Analyzer

## Slide 29: Signal Conditioning
Gain
Integration (Hardware)
AC/DC Coupling
Anti-Aliasing Filter(s)
Sample and Hold Circuit

## Slide 30: Signal Gain Circuit
X1 and X10 are Common
  Gain is simply amplification of a Signal
  Careful – Should know your vibration level and the ADC input range first!
    100mV/g accel; +-5V input range = +-50 g’s
    Can “Clip” Signal

## Slide 31: Signal Integration
Best to Integrate as close to signal source as possible
  Reduces noise

## Slide 32: AC/DC Coupling
Normally, Systems are AC coupled
  Means that there is a DC blocking Capacitor that only allows AC signal through to the system
MAARS Innovation
  DC Switch that allows AC and DC to work on the same data channel without contaminating phase
  Allows use of same channel to record data for shaft centerline (DC) and Transient data (AC)

## Slide 33: Anti-Aliasing Filters
What are they and why do I need them?
  Because “false Frequencies” are displayed when Aliasing is present in a system.
    The maximum frequency component a sampled data system can accurately handle is its Nyquist limit.
    The sample rate must be greater than or equal to two times the highest frequency component in the input signal. When this rule is violated, unwanted or undesirable signals appear in the frequency band of interest.

## Slide 34: Aliased Signals
In old western movies, as a wagon accelerates, the wheel picks up speed as expected, and then the wheel seems to slow, then stop. As the wagon further accelerates, the wheel appears to turn backwards! In reality, we know the wheel hasn't reversed because the rest of the movie action is still taking place.
What causes this phenomenon? The answer is that the shutter frame rate is not high enough to accurately capture the spinning of the wheel.

## Slide 35: Aliased Signals
False low-frequency sin wave…
  Caused by sampling too slowly
  Violated the Nyquist Criterion

## Slide 36: Anti-Aliasing Filters
What are they and why do I need them?
  Generally they are low-pass filters that do not pass frequencies above the ADC’s range.
  Here is a representation of an IDEAL filter…

## Slide 37: Real Anti-Aliasing Filters
Trade-offs: Elliptic, Chebyshev, Butterworth and Bessel
  Elliptic – sharpest rolloff, highest ripple
  Bessel – Lowest ripple, fat rolloff.
    key advantage is that it has a linear phase response

## Slide 38: Sample and Hold Circuit
Purpose is to take a snapshot of the sensor signal and hold the value.
  The ADC must have a stable signal in order to accurately perform a conversion.
  The switch connects the capacitor to the signal conditioning circuit once every sample period.
    The capacitor then holds the voltage value measured until a new sample is acquired.

## Slide 39: Data Acquisition and Storage
Analog to Digital Converter
  Hard disk vs. Flash Memory
  Physical download vs. Ethernet file Transfer
  FFT Conversion
    Windowing

## Slide 40: ADC Analog-to-Digital Converters
The purpose of the analog to digital converter is to quantize the input signal from the S&H
The input voltage can range from 0 to Vref
  What this means is that the voltage reference of the ADC is used to set the conversion range
  0V input will cause the converter to output all zeros.
  If the input to the ADC is equal to or larger than Vref, then the converter will output all ones.
  For inputs between these two voltages, the ADC will output binary numbers corresponding to the signal level.

## Slide 41: ADC Analog-to-Digital Converters
Dynamic Range
  Usually defined in dB, depends on the number of bits used by the ADC
    For example, a 12 bit ADC has 212 possible data values, or 4,096 “steps” between the lowest and highest values the ADC can see (0 to 5 Volts, typ.)
    8-bit is 256 steps
    16-bit is 65,536 steps, so more is better, right?

## Slide 42: ADC Analog-to-Digital Converters
Wrong!
Steve Goldman’s Book – pp.46-47
  “Dynamic Range: The Big Lie”
    “That the A/D Converter can sense one part in 16 binary bins is no assurance that the analog circuitry is good enough to insure that the information going into the lower bins is not contaminated by electrical noise.”

## Slide 43: ADC Analog-to-Digital Converters
Dynamic Range
  For a 12 bit ADC…20 log (4095/1) = 72 db
    Theoretical only, electronic noise reduces to 65 db
  For a 16 bit ADC…20 log (65536/1) = 96 db
    Electronic noise may make this only 80 db
Massively more data to manipulate w/o much practical gain in Dynamic Range.

## Slide 44: ADC Analog-to-Digital Converters
Sampling Rate
  “Real-Time” Rate in samples/sec
    60,000 samples per sec/2.56 = 23,437 Hz Fmax
    May also get divided by the number of channels in a multi-channel system

## Slide 45
Rotation
Heavy Spot
1 revolution
Time
Amplitude
0
+
-
Time Waveform
3600 rpm = 3600 cycles per minute
60 Hz = 60 cycles per second
1 order = one times turning speed
360 degrees

## Slide 46
1000 rpm
1 revolution
Time
Amplitude
0
+
-
Time Waveform
4 blades = vibration occurs 4 times per revolution
4 x 1000 rpm = vibration occurs at 4000 cycles per minute
= 4000 cpm

## Slide 47
12 tooth
gear
1000 rpm
1 revolution
Time
Amplitude
0
+
-
Time Waveform
12 teeth are meshing every revolution of the gear
12 x 1000 rpm = vibration occurs at 12,000 cycles per minute
= 2,000 cpm = 200 Hz

## Slide 48
Time
0
+
-
Time
0
+
Time
0
+
-
-

## Slide 49
Time
0
-
+
Time Waveform contains all the different frequencies mixed together
Complex Time Waveform

## Slide 50
Time Waveform contains all the different frequencies mixed together
Complex Time Waveform

## Slide 51
We are now entering the Frequency Domain….
FFT - Fast Fourier Transform
Separates individual frequencies
Detects how much vibration at each frequency

## Slide 52
TIME WAVEFORM
AMPLITUDE VS TIME
Notes: A typical time waveform
Horizontal axis is time in milliseconds
Vertical axis is amplitude of acceleration in G-s

## Slide 53
Amplitude
Frequency
Amplitude
Time
Amplitude
Time
Amplitude
Time
Frequency
Notes: The FFT process (Fast Fourier Transform) is a process which separates out the numerous pure frequencies (Sine Waves) that make up the Time Waveform.
After the Time Waveform is processed the discrete frequencies and their associated magnitudes are displayed.

## Slide 54
Time
0
-
Time
0
+
Time
0
+
-
-
Frequency
Frequency
Frequency
1x
4x
12x

## Slide 55
Predefined Spectrum Analysis Bands
20000
0.3
0.6
0.9
1.2
1.5
1.8
1xRPM - BALANCE
2xRPM - ALIGNMENT
3-5xRPM - LOOSENESS
5000
10000
15000
Frequency Hz
5-25xRPM
25-65xRPM
ANTI-FRICTION BEARINGS & GEARMESH

## Slide 56
Frequency Band Alarming and Trending
Alarm
Trend of
Balance
Trend of Bearings
Alarm
Amplitude
Sub-
Har
monic
1X
2X
Bearing
Bearing
Gears
Bearing
1x
2x
25-60 x
.5
in/sec
.1
in/sec
Time
(Days)
Time
(Days)
Warning
Warning
Notes: Selective frequency bands are a reliable alarming technique for determining specific machine problems
Much research has been put into the suggested levels and placement of selective frequency bands for common machinery

## Slide 57: FFT or Fast Fourier Transformer
We Use a FFT to pull data from the time waveform then input to a discrete Fourier transform (DFT), a mathematical process or algorithm that transforms a waveform into the components to a frequency spectrum of the input signal from the time waveform.

## Slide 58: Mathematical Modelling

## Slide 59: Discrete Fourier Transform (DFT)
The DFT provides uniformly spaced samples of the Discrete-Time Fourier Transform (DTFT)
DFT definition:
Requires N2 complex multiplies and N(N-1) complex additions

## Slide 60: Faster DFT computation?
Take advantage of the symmetry and periodicity of the complex exponential (let WN=e-j2p/N)
  symmetry:
  periodicity:
Note that two length N/2 DFTs take less computation than one length N DFT: 2(N/2)2<N2
Algorithms that exploit computational savings are collectively called Fast Fourier Transforms

## Slide 61: DIT Algorithm (cont.)
Result is the sum of two N/2 length DFTs
Then repeat decomposition of N/2 to N/4 DFTs, etc.
X[0…7]
x[0,2,4,6]
x[1,3,5,7]
N/2
DFT
N/2
DFT

## Slide 62: Machine Vibration
Time record
Time window
Signal from Transducer (in our case Acc. Readings)

## Slide 63: RMS vs. Peak or Peak to Peak
RMS
peak
peak to peak
RMS = 0.707 of peak

## Slide 64: Forces Causing Vibration
~99%
Rotational
~1%
Impacts
~0.1%
Friction

## Slide 65: Time domain
Frequency domain

## Slide 66: Condition Parameters - Time Domain Calculations
IMPACTS / SHOCKS
Periodic / non-periodic transients
CREST KURT SKEW
Crest Kurtosis Skewness factor
FRICTION
Periodic / non-periodic random noise
NL1 NL2 NL3 NL4
ROTATIONAL
Vibration Forces
ACCELERATION
ACC
Total RMS-value
of acceleration
VELOCITY
VEL
Total RMS-value
of velocity
Unbalance
Misalignment
Looseness
Wear
Gear damage
Coupling damage
Other shocks/ impacts
Scraping parts
Journal bearing
Used to detect and monitor

## Slide 67: Fault Frequencies Matches

## Slide 68: Analysis in the Time Waveform

## Slide 69: Analysis in the Time Waveform

## Slide 70: Analysis in the Time Waveform

## Slide 71: Windowing
Required to solve “Leakage”
  Several Types
    Uniform
    Hanning – Most Commonly used
    Hamming
    Blackman-Harris

## Slide 72: Windowing
Why do we use the Hanning Window?
  Best compromise between frequency resolution and amplitude accuracy for steady-state machinery analysis
  Uniform or Flat-Top is the best choice for transient machinery analysis.

## Slide 73: Windowing
What is leakage?
  Caused when the time waveform signal does NOT begin and end at the same point, introducing spurious frequencies.
  The Window or weighting function attenuates the signal towards the edge of the window – minimizing leakage.

## Slide 74: Windowing
Example:

## Slide 75: Windowing
Leakage Example:

## Slide 76: Windowing
Hanning Window:

## Slide 77: Types of Averaging
Linear – Most commonly used
Peak Hold – Coastdown and Impact
Exponential
  Weights most recently acquired data more heavily – used for Impact
Time Synchronous –TSA
  Triggered by tach – Shaft and Harmon.

## Slide 78: Trending Overalls
Limited Value
  Better than Nothing
  May miss some types of failures

## Slide 79: Spectral Resolution
Common Values
  100 to 3200 “Lines”
  400 or 800 typical
  Fmax/Lines = Frequency Resolution
    1000 Hz/400 lines = 2.5 Hz Resolution

## Slide 80: Spectral Integration
Where does the “Ski-Slope come from?
  Integrating Acceleration to get Velocity pops out a constant value, which is manifested as a “DC” component because it has no frequency dependence!

## Slide 81: Spectral Integration
How do we solve this problem?

## Slide 82: Spectral Integration
Truth is – we can’t!
  It’s PHYSICS!
What we can do is…
  “Zero” the first 5 or so Spectral Bins!

## Slide 83: Spectrum Analysis
Machine Component Condition
  Identified by Frequency
  Severity Indicated by Amplitude
  Rate of Deterioration Indicated by Spectral Comparison over Time

## Slide 84: Waveform Analysis
Pattern Recognition is Key
  Requires understanding of Machine Components
    Gearbox
    Bearings

## Slide 85: Waveform Analysis

## Slide 86: Orbit Analysis

## Slide 87: Machine Transients

## Slide 88: Vibration Severity
When do I make the call?
  Alarm Levels
  Fault Levels
  Do you use GM, API, ISO Guidelines?
  Risk vs. Reward

## Slide 89: Future Issues
Need to formalize the communication with other tools
  Input Details
  Output Details
  Other issues
Analysis Details
  What information will be needed ?
  How to extract that information ?
Binding Issues
  Task communication requirements

## Slide 90: Communication with tools
HW/SW Estimator
  Invocation Details
  Tool-specific output needed
Partitioner
  Bindings output and segregation of data captured.
Synthesizer tools
C / Py Function Extractor

## Slide 91: Other Issues
Profiler
  Default DataSet ?
  Machinery specific or Generic?

## Slide 92: References
1. Marne, N. S. Vibration Measurement System with Accelerometer Sensor Based on ARM [Text] / N. S. Marne, M. S. Nag-mode, R. D. Komati // International Journal of Emerging Technology and Advanced Engineering. – 2014. – Vol. 4, Issue 4. –P. 760–764.
2. Iwaniec, J. Output-Only Identification of Vibratory Machine Suspension Parameters under Exploitational Conditions [Text] /J. Iwaniec, M. Iwaniec // Solid State Phenomena. – 2016. – Vol. 248. – P. 175–185. doi: 10.4028/www.scientific.net/ssp.248.175

## Slide 93: References (contd.)
3. Kharchenko, L. Fluctuation of multi-section aboveground pipeline region under the influence of moving diagnostic pistol[Text] / L. Kharchenko, Ye. Kharchenko // Vibration in Physical Systems. – 2014. – Vol. 26. – P. 105–112.
4. Tykhan, M. New type of piezoresistive pressure sensors for environments with rapidly changing temperature [Text] / M. Ty-khan, O. Ivakhiv, V. Teslyuk // Metrology and Measurement Systems. – 2017. – Vol. 24, Issue 1. – P. 185–192. doi: 10.1515/mms-2017-0010
5. Du, W. Y. Resistive, capacitive, inductive, and magnetic sensor technologies [Text] / W. Y. Du. – Boca Raton: CRC Press,2014. – 408 p. doi: 10.1201/b17685

## Slide 94: References(Contd.)
6. Kazys, R. Ultrasonic technique for Vibration Measurements [Text] / R. Kazys, R. Sliteris, L. Mazeika // 15th World Conference on Non-Destructive Testing. – Rome, 2000.
7. Li, T. Study on the non-contact FBG vibration sensor and its application [Text] / T. Li, Y. Tan, Z. Zhou, L. Cai, S. Liu, Z. He,K. Zheng // Photonic Sensors. – 2015. – Vol. 5, Issue 2. – P. 128–136. doi: 10.1007/s13320-015-0220-9
8. Rață, G. System for Monitoring and Analysis of Vibrations at Electric Motors [Text] / G. Rață, M. Rață // Intern. Journal ofEmerging Technology and Advanced Engineering. – 2014. – Vol. XXI, Issue 3. – P. 97–104.
9. Eastern-European Journal of Enterprise Technologies ISSN 1729-3774 6/9 ( 90 ) 2017-62

## Slide 95: Appendix
Fig. i Market Review of Existing Vibrational Analyser

## Slide 96: Appendix
Fig. ii Matlab Code for FFT plots

## Slide 97: Appendix
All Codes are reposited in GitHub repo :